## Slide 1: UK NEQAS BC Altuvoct Surveys
Anna Williams

## Slide 2: Content
Survey 1- Altuvoct samples & Altuvoct-Emicizumab samples
Survey 2- Altuvoct Product Specific Calibrator samples
Notes: Requirement to supply labs with samples that could be tested with APTT reagents to show what would happen to FVIII activity results with patients on Altuvoct-Data gathering to generate information to participants
A need to supply product specific calibrator in real-world situation-can it be done for labs who use a different APTT reagent than Actin FSL?. Can Chromogenic assay kits actually be used when field studies stated that all chromogenic kits over-estimated results by 2.5 fold

## Slide 3: Survey 1-FVIII assays for post infusion monitoring
The first UK NEQAS BC Altuvoct survey
2 categories of samples;
  In-vitro spiked FVIII deficient plasma (FVIII <1IU/dL) with Altuvoct at varying concentrations including levels at 5 (EFA 24:01), 20 (EFA 24:02), 50 (EFA 24:03) & 100 (EFA 24:04) IU/dL.
  Concomitant in-vitro spiked FVIII deficient plasma with Altuvoct and Emicizumb at varying concentrations including 50IU/dL Alt & 20µg/ml EMI, 50IU/dL Alt & 50µg/ml EMI, 100IU/dL Alt & 20µg/ml EMI and 100IU/dL Alt & 50µg/ml EMI.
Participants were asked to test each sample using either a one stage or chromogenic assay using laboratory reagents and methods for measuring Altuvoct.
FVIII activity recovery submitted by participants

## Slide 4: Survey 1-One-stage assay returns
Actin FS and Synthasil were the most used one-stage APTT reagents.
Actin FS over-estimated
Synthasil underestimated the FVIII recovery of the assigned potency of efanesoctocog alfa.
The APTT reagents of Actin FSL, CK Prest, Pathromtin And Synthafax were associated with FVIII recovery close to the assigned potency of drug particularly at the 100IU/dL concentration.
Recovery ranges % for One Stage Assays (n≥10) |  |  |  |  |  | 
Sample | Actin FS | Actin FSL | CK Prest | Pathromtin | Synthafax | Synthasil
EFA 24:01 | 136-686 | 50-324 | 84-210 | 118-704 | 68-302 | 40-180
EFA 24:02 | 111.5-264 | 75-149 | 91-165 | 105-170 | 85-139.5 | 35-105
EFA 24:03 | 116.2-266 | 42.6-138.8 | 79.2-146 | 103.2-151 | 76-130.4 | 34-92
EFA 24:04 | 98.3-265.3 | 34.8-126.7 | 80-142 | 97-137 | 61.3-96.7 | 29.4-87

## Slide 5: Survey 1-Chromogenic assay returns
Chromogenic assay results, regardless of the manufacturer displayed over estimation in FVIII activity recovery compared to the potency of drug.
Recovery range % for Chromogenic Assays (n≥10) |  |  | 
Sample | Biophen FVIII | Coamatic Chromogenix | Siemens
EFA 24:01 | 210-340 | 80-334 | 78-418
EFA 24:02 | 178-420 | 65-262.5 | 105.5-262.5
EFA 24:03 | 142.6-273.4 | 74.4-420 | 101.8-272.6
EFA 24:04 | 183-240 | 85.2-406 | 97-254

## Slide 6: Survey 1-EFA & Emi results
Unmodified one stage assays substantially overestimated FVIII activity of EFA at all levels of drug in the presence of Emicizumab.
All chromogenic assay kits over-estimated the level of EFA in all samples provided to participants.

## Slide 7: Survey 1-Summary
Samples containing EFA alone
5 IU/dl, 20 IU/dl and 50 EFA there was some over estimation in one stage assays with CK Prest; Pathromtin and Synthafax.
At all concentrations the reagents unsuitable include Actin FS (over estimation) and Synthasil (under estimation) used for One stage Assays.
All types of Chromogenic assays (over estimation).
Samples containing a combination of EFA and Emicizumab
Unmodified one stage assays substantially overestimated FVIII activity of EFA at all levels of drug in the presence of Emicizumab irrespective of reagent.
Chromogenic assays with bovine FX in the kit reagents overestimate EFA. The chromogenic assay with human FX in the kit reagents (Biophen) overestimates EFA in the presence of Emicizumab to a greater extent than kits with bovine FX do.

## Slide 8: Survey 2- Product specific calibrator samples.
3 categories of samples;
    FVIII ALT CAL 25:01-Calibrator material 100IU/dL
    FVIII ALT 25:02-Post 7 days treatment level ~15IU/dL
    FVIII ALT 25:03-Post 3 days treatment level ~ 50IU/dL
Participants were asked to calibrate FVIII assays using the FVIII ALT CAL 25:01 and perform FVIII assays using their routine method and reagents on the further two post treatment samples provided (FVIII ALT 25:02 & 25:03).
FVIII activity recovery reported back to participants

## Slide 9: Survey 2-One stage assay Results
The overall median for FVIII ALT 25:02 = 15.8IU/dL and FVIII ALT 25:03 = 52.5IU/dL.
FVIII ALT 25:02 | n | median | CV | UOM | Range
Actin FSL | 17 | 14.8 | 19.5 | 0.9 | 9.0-21.1
Actin FS | 2 | 15 | NA | 0 | 15
APTT HS | 1 | 13.4 | NA | NA | NA
APTT S | 1 | 14 | NA | NA | NA
CK Prest | 3 | 17 | NA | 0.46 | 17.0-18.1
Pathromtin SL | 1 | 26.2 | NA | NA | NA
Synthasil | 12 | 16.1 | 10.9 | 0.62 | 13.0-18.0
Synthafax | 4 | 22.2 | NA | 1.63 | 18.0-24.0
Overall | 41 | 15.80 | 20.90 | 0.66 | 9.0-26.2
FVIII ALT 25:03 | n | median | CV | UOM | Range
Actin FSL | 17 | 49 | 8 | 1.21 | 43.7-56.0
Actin FS | 2 | 50 | NA | 2.5 | 48.0-52.0
APTT HS | 1 | 39.7 | NA | NA | NA
APTT S | 1 | 51 | NA | NA | NA
CK Prest | 3 | 56 | NA | 1.94 | 54.9-60.0
Pathromtin SL | 1 | 79.4 | NA | NA | NA
Synthasil | 12 | 53.1 | 10.5 | 1.95 | 41.5
Synthafax | 4 | 56.6 | NA | 3.9 | 51.0-64.9
Overall | 41 | 52.5 | 13.4 | 1.4 | 39.7

## Slide 10: Survey 2-Chromogenic assay Results
The overall median for FVIII ALT 25:02 = 14.2 IU/dL and FVIII ALT 25:03 = 49.0 IU/dL.
FVIII ALT 25:02 | n | median | CV | UOM | Range
Siemens | 11 | 14 | NA | 0.3 | 12.0-15.0
Electrochrome | 1 | 14.2 | NA | NA | NA
Chromogenix coamatic | 3 | 17 | NA | 1.01 | 14.6-17.0
Hyphen biophen- human | 3 | 16 | NA | 1.13 | 14.9-18.0
Hyphen biophen human/bovine mix (Varient) | 1 | 13.6 | NA | NA | NA
Overall | 19 | 14.2 | 10.3 | 0.43 | 12.0-18.0
FVIII ALT 25:03 | n | median | CV | UOM | Range
Siemens | 11 | 48.4 | 10.1 | 1.81 | 34.0-53.5
Electrochrome | 1 | 60.6 | NA | NA | NA
Chromogenix coamatic | 3 | 56 | NA | 4.8 | 56.0-67.6
Hyphen biophen- human | 3 | 52.3 | NA | 10.2 | 34.0-53.5
Hyphen biophen human/bovine mix (Varient) | 1 | 48.2 | NA | NA | NA
Overall | 19 | 49 | 16.1 | 2.4 | 34.0-67.6

## Slide 11: Survey 2-Conclusion
Product specific calibrator material confirms that a larger range of APTT reagents can be used when measuring samples containing Altuvoct.
Including the widely used APTT reagents of Synthasil and Actin FS, that previously were suggested to over- and under-estimate FVIII Activity compared to target potency of sample1&2.
However, this is still not true for the APTT reagents of Pathromtin SL and Synthafax , as both over-estimate FVIII results by 1.5 fold.
Documented literature and our UK NEQAS BC global study suggested that chromogenic assays should never be used when monitoring samples containing Altuvoct. This survey has shown that with the use of a product specific calibrator, any chromogenic assay can be used and the FVIII activity recovery is acceptable towards the target potency.

## Slide 12: Thank you for listening anna.williams79@nhs.net